LDI inversion from WUSTL (v1.0)

The m-files needed for LDI are located in the LDI_fitting folder. Add this folder (with subfolders) to your matlab path.

There are three test cases

- 2D spiral data from a cylindrical gel phantom filled with gelatin/glycerol blend. This data was acquired on the 1.5T Avanto with paddle actuator in Sept 2013 at 50 Hz with two motion encoding gradient cycles.

- 2D EPI  MRE data at 50 Hz from a cylindrical phantom filled with Sylgard 529 (1:1.25 blend). It is stiffer than gelatin and more viscoelastic. One set of data was acquired at a motion encoding gradient strength of 10 mT/m, the second set was acquired 25.9mT/m. We used the pillow actuator and the motion is primarily in the y and z directions.

- 2D EPI MRE data at 50 Hz from a human subject at 50 Hz. Similar to the Sylgard phantom,  there are two datasets, one for 10 mT/m and one for 25.9 mT/m.

Each test case has a mat file with the following variables:

The stored variables are:

- PCdata  (  80 x 80 x 24 x 8 x 3):  Phase Contrast  data, masked and  unwrapped

  - X (scanner RL), Y (scanner AP) and Z (scanner FH)

Displacements stored in radians

- MAG  Magnitude images (not masked), corresponding to PC data

- mask binary mask and nanmask (binary mask with NaNs replacing zeros)

- subjID

  - name, date and frequency of data set, i.e. ‘MPxxxx-201ymmdd-xxHz’

- PC2micron conversion of displacement (in radians) to microns. The value depends on sequence, motion-encoding gradient strength and frequency.

- voxelsize

  - All Trio spiral scans acquired w/ 2.0 mm isotropic voxels, and all Trio EPI MRE scans acquired w/ 3.0 mm isotropic voxels.

- userow and usecol vectors used to select portion of image slices to display. To display entire image volume, set these to 1:size(PCdata,1) and 1:size(PCdata,2) respectively.

- unwrapflag,unwrapmethod used to track whether data had to be unwrapped. If unwrapflag = 0, no unwrapping was necessary.

- reorder_flag, for EP MRE data, indicates that time dimension has been reordered

The main script for running LDI is called run_MRE_LDI_script_McGarry.m. The script has seven steps that can be run in sections using the editor or from start to finish.

There are various times at which plots are made with a function called vis5d_profile.m. This function plots a 3D, 4D or 5D array

## Step 1: Load Data

This allows you to choose a file and specify any needed parameters that are not in your input file.

## Step 2:  Bulk Motion removal

Bulk motion is removed if desired.  To remove bulk motion, call the following function:

Another variable, called PCdata_BMR is created  along with coefficients describing the bulk motion at each time point. PCdata_BMR has same dimensions as PCdata but with Bulk Motion Removed (BMR). To see bulk motion, subtract PCdata_BMR from PC_data

Note that removing bulk motion does not affect LDI calculations if curl is used as the input

## Step 3: Select slices for fitting

Plots data from step 2, checks for PC2micron and allows user to select slices for fitting.

## Step 4:  Calculate strains, curl, OSS_SNR

Calculate strains and curl using Matt McGarry’s program for OSS_SNR calculation. Displays table of strains (10%, 50% (median), 90% values for each component).

Step 5: Plot smoothed displacement and curl fields if desired.

## Step 6:  Calculate displacement and curl for LDI

You should create  a sub-folder in which to store LDI fitting inputs and results, e.g. “LDI_fitting”

The script will create a file

<subjID>-<todays date>-curl_LDI_fitting_inputs.mat that stores data for LDI fitting.

## Step 7: Run  LDI

User selects fitting parameters (or accepts default values):

Delta – estimate of the noisiness of the Laplacian, usually use 0.01 or 0.0001

fitkernel – size of fitting region, must be 3, 5, 7 etc.  I usually vary this parameter.

smoothfilter – two element vector [sd kernelsize ]  specifying Gaussian filter parameters used to smooth the input data (displacement or curl). I usually use  [1 3] or [1 5].

The parameter nanthresh is currently fixed at 0.5 This parameter specifies the percentage of voxels in a fitting kernel that can be NaN before the point is assigned the NaN value. The value of 0.5 seems to work well. A lower value will cause more points to be discarded near the boundary of the mask, a higher value may cause some spurious fitting values to be included.

After the parameters are set script calls the function:

run_LDI_fun(subjID,usecurl,fitkernel,delta,smoothfilter,nanthresh,LDIdir) which performs the fits, then stores and plots the results. Summary statistics of the real part of the shear modulus, imaginary part of the shear modulus, and the normalized residual error are shown.